## Munka1
Model | Brand name | Medical device (Product name) | Basic UDI-DI | Reference number of Technical documentation (TF) | MD intended purpose | Indications for use as per IFU | EMDN code | GMDN code | MDA / MDN code of MD | MDS horizontal codes of MD | MDT horizontal codes of MD | Classification - MDD | Classification - MDR | All applicable classification rules | Justification of the classification | Is the device a new development to the market? (Yes / No) | Is the product supplied sterile? (Yes / No) | Sterilisation method (Click to select, if applicable supplement instead of "specify" ) | Location of sterilisation (Click to select) | Is the sterilisation process of the specified products validated? (Click to select, if applicable, please attach a copy of the sterilization validation report.) | Is the device wholly or mainly absorbed? (Yes / No) | Is the device locally dispersed in the human body or is intended to undergo a chemical change in the body? (Yes / No) | Is the device in systems or procedure packs? (Yes / No) | Please enclose product information / brochures / instructions for use including detailed description of intended use of a product. List the attached documents (Name, date, revision number) | Specification of the conformity assessment procedure according to EU Regulation 2017/745 (Click to select) | If you sell products under your own company name, which are produced by other companies (art. 16 of Regulation (EU) 2017/745), did the original equipment manufacturer (OEM) already carry out a conformity assessment procedure? Please enclose copies of already existing OEM EU certificates (if the products were assessed by other NB). | Critical Supplier to the MD and address, Outsourced process (e.g. sterilisation, production, etc.) | Notified Body - MDD | Notified Body - MDR
640PM | Q-Flex Trifocal Preloaded | Posterior-chamber intraocular lens, pseudophakic, hydrophilic | 599302640PMYV | TF_01_B | Medicontur hydrophilic posterior chamber IOLs are intended for implantation into the capsular bag in the posterior chamber of the eye to replace the human crystalline lens in adult patients to improve vision (either distance vision or distance, intermediate, near vision; and additionally, optionally reduce the cylindrical power of the eye). | Medicontur hydrophilic posterior chamber trifocal IOLs are indicated to improve vision at far, intermediate and near distances in adults with cataract and/or ametropia (hyperopia, myopia) and/or presbyopia, secondary after removal of the crystalline lens. | P030102100202– IOLS, APHAKIC, MULTIFOCAL, ASPHERIC, HYDROPHILIC, ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Hydrophilic IOLs are implantable devices and they: - are not intended to be placed in the teeth; - are not intended to be used in direct contact with the heart, the central circulatory system or the central nervous system; - do not have a biological effect or are not wholly or mainly absorbed; - are not intended to undergo chemical change in the body; - are not intended to administer medicinal products. - are not active implantable devices or their accessories; - are not breast implants or surgical meshes; - are not total or partial joint replacements; - are not spinal disc replacement implants, or, are not implantable devices that come into contact with the spinal column. | No | Yes | Moist heat | In-house | Yes | No | No | Yes | LB-005-5100-00-V06 HL Preloaded Flex IFU EN | Annex IX | N/A | Benz R&D, USA, Sarasota 6447 Parkland Drive, FL34230; Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
677AD | Bi-Flex | Posterior-chamber intraocular lens, pseudophakic, hydrophilic | 599302677ADYN | TF_01_B | Medicontur hydrophilic posterior chamber IOLs are intended for implantation into the capsular bag in the posterior chamber of the eye to replace the human crystalline lens in adult patients to improve vision (either distance vision or distance, intermediate, near vision; and additionally, optionally reduce the cylindrical power of the eye). | Medicontur hydrophilic posterior chamber monofocal IOLs are indicated to improve vision at far distance in adults with cataract and/or ametropia (hyperopia, myopia), secondary to removal of the crystalline lens. | P030102090202 – IOLS, APHAKIC, MONOFOCAL, ASPHERIC, HYDROPHILIC, ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Same as above | No | Yes | Moist heat | In-house | Yes | No | No | No | LB-003-5101-00-V06 POSTERIOR CHAMBER HL IOL IFU EN | Annex IX | N/A | Benz R&D, USA, Sarasota 6447 Parkland Drive, FL34230; Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
677ADY | Bi-Flex | Posterior-chamber intraocular lens, pseudophakic, hydrophilic, blue light filter | 599302677ADYRD | TF_01_B | Medicontur hydrophilic posterior chamber IOLs are intended for implantation into the capsular bag in the posterior chamber of the eye to replace the human crystalline lens in adult patients to improve vision (either distance vision or distance, intermediate, near vision; and additionally, optionally reduce the cylindrical power of the eye). | Medicontur hydrophilic posterior chamber monofocal IOLs are indicated to improve vision at far distance in adults with cataract and/or ametropia (hyperopia, myopia), secondary to removal of the crystalline lens. | P030102090202 – IOLS, APHAKIC, MONOFOCAL, ASPHERIC, HYDROPHILIC, ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Same as above | No | Yes | Moist heat | In-house | Yes | No | No | No | LB-003-5101-00-V06 POSTERIOR CHAMBER HL IOL IFU EN | Annex IX | N/A | Benz R&D, USA, Sarasota 6447 Parkland Drive, FL34230; Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
677CMTY | Trifocal Toric Bi-Flex Liberty Preloaded | Posterior-chamber intraocular lens, pseudophakic, hydrophilic, blue light filter | 599302677CMTYP3 | TF_01_B | Medicontur hydrophilic posterior chamber IOLs are intended for implantation into the capsular bag in the posterior chamber of the eye to replace the human crystalline lens in adult patients to improve vision (either distance vision or distance, intermediate, near vision; and additionally, optionally reduce the cylindrical power of the eye). | Medicontur hydrophilic posterior chamber trifocal toric IOLs are indicated to improve vision at far, intermediate and near distances and to reduce the cylindrical power of the eye in adults with cataract and/or ametropia (hyperopia, myopia, astigmatism), secondary to removal of the crystalline lens. | P030102100302– IOLs, APHAKIC, MULTIFOCAL, TORIC, HYDROPHILIC ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Same as above | No | Yes | Moist heat | In-house | Yes | No | No | Yes | LB-005-5100-00-V06 HL Preloaded Flex IFU EN | Annex IX | N/A | Benz R&D, USA, Sarasota 6447 Parkland Drive, FL34230; Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
677CMY | Trifocal Bi-Flex Liberty Preloaded | Posterior-chamber intraocular lens, pseudophakic, hydrophilic, blue light filter | 599302677CMYSJ | TF_01_B | Medicontur hydrophilic posterior chamber IOLs are intended for implantation into the capsular bag in the posterior chamber of the eye to replace the human crystalline lens in adult patients to improve vision (either distance vision or distance, intermediate, near vision; and additionally, optionally reduce the cylindrical power of the eye). | Medicontur hydrophilic posterior chamber trifocal IOLs are indicated to improve vision at far, intermediate and near distances in adults with cataract and/or ametropia (hyperopia, myopia) and/or presbyopia, secondary after removal of the crystalline lens. | P030102100202– IOLS, APHAKIC, MULTIFOCAL, ASPHERIC, HYDROPHILIC, ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Same as above | No | Yes | Moist heat | In-house | Yes | No | No | Yes | LB-005-5100-00-V06 HL Preloaded Flex IFU EN | Annex IX | N/A | Benz R&D, USA, Sarasota 6447 Parkland Drive, FL34230; Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
677CTAY | Bi-Flex T Preloaded | Posterior-chamber intraocular lens, pseudophakic, hydrophilic, blue light filter | 599302677CTAYND | TF_01_B | Medicontur hydrophilic posterior chamber IOLs are intended for implantation into the capsular bag in the posterior chamber of the eye to replace the human crystalline lens in adult patients to improve vision (either distance vision or distance, intermediate, near vision; and additionally, optionally reduce the cylindrical power of the eye). | Medicontur hydrophilic posterior chamber monofocal toric IOLs are indicated to improve vision at far distance and to reduce the cylindrical power of the eye in adults with cataract and/or ametropia (hyperopia, myopia, astigmatism), secondary to removal of the crystalline lens. | P030102090302 – IOLS, APHAKIC, MONOFOCAL, TORIC, ASPHERIC, HYDROPHILIC, ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Same as above | No | Yes | Moist heat | In-house | Yes | No | No | Yes | LB-005-5100-00-V06 HL Preloaded Flex IFU EN | Annex IX | N/A | Benz R&D, USA, Sarasota 6447 Parkland Drive, FL34230; Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
677MTY | Trifocal Toric Bi-Flex Liberty | Posterior-chamber intraocular lens, pseudophakic, hydrophilic, blue light filter | 599302677MTYUR | TF_01_B | Medicontur hydrophilic posterior chamber IOLs are intended for implantation into the capsular bag in the posterior chamber of the eye to replace the human crystalline lens in adult patients to improve vision (either distance vision or distance, intermediate, near vision; and additionally, optionally reduce the cylindrical power of the eye). | Medicontur hydrophilic posterior chamber trifocal toric IOLs are indicated to improve vision at far, intermediate and near distances and to reduce the cylindrical power of the eye in adults with cataract and/or ametropia (hyperopia, myopia, astigmatism), secondary to removal of the crystalline lens. | P030102100302– IOLs, APHAKIC, MULTIFOCAL, TORIC, HYDROPHILIC ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Same as above | No | Yes | Moist heat | In-house | Yes | No | No | No | LB-003-5101-00-V06 POSTERIOR CHAMBER HL IOL IFU EN | Annex IX | N/A | Benz R&D, USA, Sarasota 6447 Parkland Drive, FL34230; Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
677MY | Trifocal Bi-Flex Liberty | Posterior-chamber intraocular lens, pseudophakic, hydrophilic, blue light filter | 599302677MY37 | TF_01_B | Medicontur hydrophilic posterior chamber IOLs are intended for implantation into the capsular bag in the posterior chamber of the eye to replace the human crystalline lens in adult patients to improve vision (either distance vision or distance, intermediate, near vision; and additionally, optionally reduce the cylindrical power of the eye). | Medicontur hydrophilic posterior chamber trifocal IOLs are indicated to improve vision at far, intermediate and near distances in adults with cataract and/or ametropia (hyperopia, myopia) and/or presbyopia, secondary after removal of the crystalline lens. | P030102100202– IOLS, APHAKIC, MULTIFOCAL, ASPHERIC, HYDROPHILIC, ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Same as above | No | Yes | Moist heat | In-house | Yes | No | No | No | LB-003-5101-00-V06 POSTERIOR CHAMBER HL IOL IFU EN | Annex IX | N/A | Benz R&D, USA, Sarasota 6447 Parkland Drive, FL34230; Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
677P | Bi-Flex Preloaded | Posterior-chamber intraocular lens, pseudophakic, hydrophilic | 599302677PE9 | TF_01_B | Medicontur hydrophilic posterior chamber IOLs are intended for implantation into the capsular bag in the posterior chamber of the eye to replace the human crystalline lens in adult patients to improve vision (either distance vision or distance, intermediate, near vision; and additionally, optionally reduce the cylindrical power of the eye). | Medicontur hydrophilic posterior chamber monofocal IOLs are indicated to improve vision at far distance in adults with cataract and/or ametropia (hyperopia, myopia), secondary to removal of the crystalline lens. | P030102090202 – IOLS, APHAKIC, MONOFOCAL, ASPHERIC, HYDROPHILIC, ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Same as above | No | Yes | Moist heat | In-house | Yes | No | No | Yes | LB-005-5100-00-V06 HL Preloaded Flex IFU EN | Annex IX | N/A | Benz R&D, USA, Sarasota 6447 Parkland Drive, FL34230; Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
677PMY | Trifocal Bi-Flex Liberty Preloaded | Posterior-chamber intraocular lens, pseudophakic, hydrophilic, blue light filter | 599302677PMYUK | TF_01_B | Medicontur hydrophilic posterior chamber IOLs are intended for implantation into the capsular bag in the posterior chamber of the eye to replace the human crystalline lens in adult patients to improve vision (either distance vision or distance, intermediate, near vision; and additionally, optionally reduce the cylindrical power of the eye). | Medicontur hydrophilic posterior chamber trifocal IOLs are indicated to improve vision at far, intermediate and near distances in adults with cataract and/or ametropia (hyperopia, myopia) and/or presbyopia, secondary after removal of the crystalline lens. | P030102100202– IOLS, APHAKIC, MULTIFOCAL, ASPHERIC, HYDROPHILIC, ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Same as above | No | Yes | Moist heat | In-house | Yes | No | No | Yes | LB-005-5100-00-V06 HL Preloaded Flex IFU EN | Annex IX | N/A | Benz R&D, USA, Sarasota 6447 Parkland Drive, FL34230; Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
677PY | Bi-Flex Preloaded | Posterior-chamber intraocular lens, pseudophakic, hydrophilic, blue light filter | 599302677PY3G | TF_01_B | Medicontur hydrophilic posterior chamber IOLs are intended for implantation into the capsular bag in the posterior chamber of the eye to replace the human crystalline lens in adult patients to improve vision (either distance vision or distance, intermediate, near vision; and additionally, optionally reduce the cylindrical power of the eye). | Medicontur hydrophilic posterior chamber monofocal IOLs are indicated to improve vision at far distance in adults with cataract and/or ametropia (hyperopia, myopia), secondary to removal of the crystalline lens. | P030102090202 – IOLS, APHAKIC, MONOFOCAL, ASPHERIC, HYDROPHILIC, ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Same as above | No | Yes | Moist heat | In-house | Yes | No | No | Yes | LB-005-5100-00-V06 HL Preloaded Flex IFU EN | Annex IX | N/A | Benz R&D, USA, Sarasota 6447 Parkland Drive, FL34230; Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
677TAY | Bi-Flex T | Posterior-chamber intraocular lens, pseudophakic, hydrophilic, blue light filter | 599302677TAYU3 | TF_01_B | Medicontur hydrophilic posterior chamber IOLs are intended for implantation into the capsular bag in the posterior chamber of the eye to replace the human crystalline lens in adult patients to improve vision (either distance vision or distance, intermediate, near vision; and additionally, optionally reduce the cylindrical power of the eye). | Medicontur hydrophilic posterior chamber monofocal toric IOLs are indicated to improve vision at far distance and to reduce the cylindrical power of the eye in adults with cataract and/or ametropia (hyperopia, myopia, astigmatism), secondary to removal of the crystalline lens. | P030102090302 – IOLS, APHAKIC, MONOFOCAL, TORIC, ASPHERIC, HYDROPHILIC, ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Same as above | No | Yes | Moist heat | In-house | Yes | No | No | No | LB-003-5101-00-V06 POSTERIOR CHAMBER HL IOL IFU EN | Annex IX | N/A | Benz R&D, USA, Sarasota 6447 Parkland Drive, FL34230; Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
690AD | Z-Flex | Posterior-chamber intraocular lens, pseudophakic, hydrophilic | 599302690ADXZ | TF_01_B | Medicontur hydrophilic posterior chamber IOLs are intended for implantation into the capsular bag in the posterior chamber of the eye to replace the human crystalline lens in adult patients to improve vision (either distance vision or distance, intermediate, near vision; and additionally, optionally reduce the cylindrical power of the eye). | Medicontur hydrophilic posterior chamber monofocal IOLs are indicated to improve vision at far distance in adults with cataract and/or ametropia (hyperopia, myopia), secondary to removal of the crystalline lens. | P030102090202 – IOLS, APHAKIC, MONOFOCAL, ASPHERIC, HYDROPHILIC, ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Same as above | No | Yes | Moist heat | In-house | Yes | No | No | No | LB-003-5101-00-V06 POSTERIOR CHAMBER HL IOL IFU EN | Annex IX | N/A | Benz R&D, USA, Sarasota 6447 Parkland Drive, FL34230; Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
690ADY | Z-Flex | Posterior-chamber intraocular lens, pseudophakic, hydrophilic, blue light filter | 599302690ADYQJ | TF_01_B | Medicontur hydrophilic posterior chamber IOLs are intended for implantation into the capsular bag in the posterior chamber of the eye to replace the human crystalline lens in adult patients to improve vision (either distance vision or distance, intermediate, near vision; and additionally, optionally reduce the cylindrical power of the eye). | Medicontur hydrophilic posterior chamber monofocal IOLs are indicated to improve vision at far distance in adults with cataract and/or ametropia (hyperopia, myopia), secondary to removal of the crystalline lens. | P030102090202 – IOLS, APHAKIC, MONOFOCAL, ASPHERIC, HYDROPHILIC, ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Same as above | No | Yes | Moist heat | In-house | Yes | No | No | No | LB-003-5101-00-V06 POSTERIOR CHAMBER HL IOL IFU EN | Annex IX | N/A | Benz R&D, USA, Sarasota 6447 Parkland Drive, FL34230; Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
690CMTY | Trifocal Toric Z-Flex Preloaded | Posterior-chamber intraocular lens, pseudophakic, hydrophilic, blue light filter | 599302690CMTYMG | TF_01_B | Medicontur hydrophilic posterior chamber IOLs are intended for implantation into the capsular bag in the posterior chamber of the eye to replace the human crystalline lens in adult patients to improve vision (either distance vision or distance, intermediate, near vision; and additionally, optionally reduce the cylindrical power of the eye). | Medicontur hydrophilic posterior chamber trifocal toric IOLs are indicated to improve vision at far, intermediate and near distances and to reduce the cylindrical power of the eye in adults with cataract and/or ametropia (hyperopia, myopia, astigmatism), secondary to removal of the crystalline lens. | P030102100302– IOLs, APHAKIC, MULTIFOCAL, TORIC, HYDROPHILIC ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Same as above | No | Yes | Moist heat | In-house | Yes | No | No | Yes | LB-005-5100-00-V06 HL Preloaded Flex IFU EN | Annex IX | N/A | Benz R&D, USA, Sarasota 6447 Parkland Drive, FL34230; Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
690CMY | Trifocal Z-Flex Preloaded | Posterior-chamber intraocular lens, pseudophakic, hydrophilic, blue light filter | 599302690CMYRP | TF_01_B | Medicontur hydrophilic posterior chamber IOLs are intended for implantation into the capsular bag in the posterior chamber of the eye to replace the human crystalline lens in adult patients to improve vision (either distance vision or distance, intermediate, near vision; and additionally, optionally reduce the cylindrical power of the eye). | Medicontur hydrophilic posterior chamber trifocal IOLs are indicated to improve vision at far, intermediate and near distances in adults with cataract and/or ametropia (hyperopia, myopia) and/or presbyopia, secondary after removal of the crystalline lens. | P030102100202– IOLS, APHAKIC, MULTIFOCAL, ASPHERIC, HYDROPHILIC, ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Same as above | No | Yes | Moist heat | In-house | Yes | No | No | Yes | LB-005-5100-00-V06 HL Preloaded Flex IFU EN | Annex IX | N/A | Benz R&D, USA, Sarasota 6447 Parkland Drive, FL34230; Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
690CTAY | Z-Flex T Preloaded | Posterior-chamber intraocular lens, pseudophakic, hydrophilic, blue light filter | 599302690CTAYLS | TF_01_B | Medicontur hydrophilic posterior chamber IOLs are intended for implantation into the capsular bag in the posterior chamber of the eye to replace the human crystalline lens in adult patients to improve vision (either distance vision or distance, intermediate, near vision; and additionally, optionally reduce the cylindrical power of the eye). | Medicontur hydrophilic posterior chamber monofocal toric IOLs are indicated to improve vision at far distance and to reduce the cylindrical power of the eye in adults with cataract and/or ametropia (hyperopia, myopia, astigmatism), secondary to removal of the crystalline lens. | P030102090302 – IOLS, APHAKIC, MONOFOCAL, TORIC, ASPHERIC, HYDROPHILIC, ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Same as above | No | Yes | Moist heat | In-house | Yes | No | No | Yes | LB-005-5100-00-V06 HL Preloaded Flex IFU EN | Annex IX | N/A | Benz R&D, USA, Sarasota 6447 Parkland Drive, FL34230; Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
690MTY | Z-Flex Trifocal Toric | Posterior-chamber intraocular lens, pseudophakic, hydrophilic, blue light filter | 599302690MTYTW | TF_01_B | Medicontur hydrophilic posterior chamber IOLs are intended for implantation into the capsular bag in the posterior chamber of the eye to replace the human crystalline lens in adult patients to improve vision (either distance vision or distance, intermediate, near vision; and additionally, optionally reduce the cylindrical power of the eye). | Medicontur hydrophilic posterior chamber trifocal toric IOLs are indicated to improve vision at far, intermediate and near distances and to reduce the cylindrical power of the eye in adults with cataract and/or ametropia (hyperopia, myopia, astigmatism), secondary to removal of the crystalline lens. | P030102100302– IOLs, APHAKIC, MULTIFOCAL, TORIC, HYDROPHILIC ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Same as above | No | Yes | Moist heat | In-house | Yes | No | No | No | LB-003-5101-00-V06 POSTERIOR CHAMBER HL IOL IFU EN | Annex IX | N/A | Benz R&D, USA, Sarasota 6447 Parkland Drive, FL34230; Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
690MY | Z-Flex Trifocal | Posterior-chamber intraocular lens, pseudophakic, hydrophilic, blue light filter | 599302690MY2J | TF_01_B | Medicontur hydrophilic posterior chamber IOLs are intended for implantation into the capsular bag in the posterior chamber of the eye to replace the human crystalline lens in adult patients to improve vision (either distance vision or distance, intermediate, near vision; and additionally, optionally reduce the cylindrical power of the eye). | Medicontur hydrophilic posterior chamber trifocal IOLs are indicated to improve vision at far, intermediate and near distances in adults with cataract and/or ametropia (hyperopia, myopia) and/or presbyopia, secondary after removal of the crystalline lens. | P030102100202– IOLS, APHAKIC, MULTIFOCAL, ASPHERIC, HYDROPHILIC, ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Same as above | No | Yes | Moist heat | In-house | Yes | No | No | No | LB-003-5101-00-V06 POSTERIOR CHAMBER HL IOL IFU EN | Annex IX | N/A | Benz R&D, USA, Sarasota 6447 Parkland Drive, FL34230; Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
690TAY | Z-Flex T | Posterior-chamber intraocular lens, pseudophakic, hydrophilic, blue light filter | 599302690TAYT8 | TF_01_B | Medicontur hydrophilic posterior chamber IOLs are intended for implantation into the capsular bag in the posterior chamber of the eye to replace the human crystalline lens in adult patients to improve vision (either distance vision or distance, intermediate, near vision; and additionally, optionally reduce the cylindrical power of the eye). | Medicontur hydrophilic posterior chamber monofocal toric IOLs are indicated to improve vision at far distance and to reduce the cylindrical power of the eye in adults with cataract and/or ametropia (hyperopia, myopia, astigmatism), secondary to removal of the crystalline lens. | P030102090302 – IOLS, APHAKIC, MONOFOCAL, TORIC, ASPHERIC, HYDROPHILIC, ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Same as above | No | Yes | Moist heat | In-house | Yes | No | No | No | LB-003-5101-00-V06 POSTERIOR CHAMBER HL IOL IFU EN | Annex IX | N/A | Benz R&D, USA, Sarasota 6447 Parkland Drive, FL34230; Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
640AD | Q-Flex | Posterior-chamber intraocular lens, pseudophakic, hydrophilic | 599302640ADWW | TF_01_B | Medicontur hydrophilic posterior chamber IOLs are intended for implantation into the capsular bag in the posterior chamber of the eye to replace the human crystalline lens in adult patients to improve vision (either distance vision or distance, intermediate, near vision; and additionally, optionally reduce the cylindrical power of the eye). | Medicontur hydrophilic posterior chamber monofocal IOLs are indicated to improve vision at far distance in adults with cataract and/or ametropia (hyperopia, myopia), secondary to removal of the crystalline lens. | P030102090202 – IOLS, APHAKIC, MONOFOCAL, ASPHERIC, HYDROPHILIC, ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Same as above | No | Yes | Moist heat | In-house | Yes | No | No | No | not available yet | Annex IX | N/A | Benz R&D, USA, Sarasota 6447 Parkland Drive, FL34230; Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
640ADY | Q-Flex | Posterior-chamber intraocular lens, pseudophakic, hydrophilic, blue light filter | 599302640ADYNT | TF_01_B | Medicontur hydrophilic posterior chamber IOLs are intended for implantation into the capsular bag in the posterior chamber of the eye to replace the human crystalline lens in adult patients to improve vision (either distance vision or distance, intermediate, near vision; and additionally, optionally reduce the cylindrical power of the eye). | Medicontur hydrophilic posterior chamber monofocal IOLs are indicated to improve vision at far distance in adults with cataract and/or ametropia (hyperopia, myopia), secondary to removal of the crystalline lens. | P030102090202 – IOLS, APHAKIC, MONOFOCAL, ASPHERIC, HYDROPHILIC, ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Same as above | No | Yes | Moist heat | In-house | Yes | No | No | No | not available yet | Annex IX | N/A | Benz R&D, USA, Sarasota 6447 Parkland Drive, FL34230; Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
640CMY | Trifocal Q-Flex Preloaded | Posterior-chamber intraocular lens, pseudophakic, hydrophilic, blue light filter | 599302640CMYPY | TF_01_B | Medicontur hydrophilic posterior chamber IOLs are intended for implantation into the capsular bag in the posterior chamber of the eye to replace the human crystalline lens in adult patients to improve vision (either distance vision or distance, intermediate, near vision; and additionally, optionally reduce the cylindrical power of the eye). | Medicontur hydrophilic posterior chamber trifocal IOLs are indicated to improve vision at far, intermediate and near distances in adults with cataract and/or ametropia (hyperopia, myopia) and/or presbyopia, secondary after removal of the crystalline lens. | P030102100202– IOLS, APHAKIC, MULTIFOCAL, ASPHERIC, HYDROPHILIC, ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Same as above | No | Yes | Moist heat | In-house | Yes | No | No | Yes | not available yet | Annex IX | N/A | Benz R&D, USA, Sarasota 6447 Parkland Drive, FL34230; Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
640MY | Q-Flex Trifocal | Posterior-chamber intraocular lens, pseudophakic, hydrophilic, blue light filter | 599302640MYZC | TF_01_B | Medicontur hydrophilic posterior chamber IOLs are intended for implantation into the capsular bag in the posterior chamber of the eye to replace the human crystalline lens in adult patients to improve vision (either distance vision or distance, intermediate, near vision; and additionally, optionally reduce the cylindrical power of the eye). | Medicontur hydrophilic posterior chamber trifocal IOLs are indicated to improve vision at far, intermediate and near distances in adults with cataract and/or ametropia (hyperopia, myopia) and/or presbyopia, secondary after removal of the crystalline lens. | P030102100202– IOLS, APHAKIC, MULTIFOCAL, ASPHERIC, HYDROPHILIC, ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Same as above | No | Yes | Moist heat | In-house | Yes | No | No | No | not available yet | Annex IX | N/A | Benz R&D, USA, Sarasota 6447 Parkland Drive, FL34230; Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
640P | Q-Flex Preloaded | Posterior-chamber intraocular lens, pseudophakic, hydrophilic | 599302640PD5 | TF_01_B | Medicontur hydrophilic posterior chamber IOLs are intended for implantation into the capsular bag in the posterior chamber of the eye to replace the human crystalline lens in adult patients to improve vision (either distance vision or distance, intermediate, near vision; and additionally, optionally reduce the cylindrical power of the eye). | Medicontur hydrophilic posterior chamber monofocal IOLs are indicated to improve vision at far distance in adults with cataract and/or ametropia (hyperopia, myopia), secondary to removal of the crystalline lens. | P030102090202 – IOLS, APHAKIC, MONOFOCAL, ASPHERIC, HYDROPHILIC, ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Same as above | No | Yes | Moist heat | In-house | Yes | No | No | Yes | not available yet | Annex IX | N/A | Benz R&D, USA, Sarasota 6447 Parkland Drive, FL34230; Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
640PY | Q-Flex Preloaded | Posterior-chamber intraocular lens, pseudophakic, hydrophilic, blue light filter | 599302640PYZM | TF_01_B | Medicontur hydrophilic posterior chamber IOLs are intended for implantation into the capsular bag in the posterior chamber of the eye to replace the human crystalline lens in adult patients to improve vision (either distance vision or distance, intermediate, near vision; and additionally, optionally reduce the cylindrical power of the eye). | Medicontur hydrophilic posterior chamber monofocal IOLs are indicated to improve vision at far distance in adults with cataract and/or ametropia (hyperopia, myopia), secondary to removal of the crystalline lens. | P030102090202 – IOLS, APHAKIC, MONOFOCAL, ASPHERIC, HYDROPHILIC, ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Same as above | No | Yes | Moist heat | In-house | Yes | No | No | Yes | not available yet | Annex IX | N/A | Benz R&D, USA, Sarasota 6447 Parkland Drive, FL34230; Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
677CTA | Bi-Flex T preloaded | Posterior-chamber intraocular lens, pseudophakic, hydrophilic | 599302677CTARP | TF_01_B | Medicontur hydrophilic posterior chamber IOLs are intended for implantation into the capsular bag in the posterior chamber of the eye to replace the human crystalline lens in adult patients to improve vision (either distance vision or distance, intermediate, near vision; and additionally, optionally reduce the cylindrical power of the eye). | Medicontur hydrophilic posterior chamber monofocal toric IOLs are indicated to improve vision at far distance and to reduce the cylindrical power of the eye in adults with cataract and/or ametropia (hyperopia, myopia, astigmatism), secondary to removal of the crystalline lens. | P030102090302 – IOLS, APHAKIC, MONOFOCAL, TORIC, ASPHERIC, HYDROPHILIC, ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Same as above | No | Yes | Moist heat | In-house | Yes | No | No | Yes | not available yet | Annex IX | N/A | Benz R&D, USA, Sarasota 6447 Parkland Drive, FL34230; Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
677M | Trifocal Bi-Flex Liberty | Posterior-chamber intraocular lens, pseudophakic, hydrophilic | 599302677ME3 | TF_01_B | Medicontur hydrophilic posterior chamber IOLs are intended for implantation into the capsular bag in the posterior chamber of the eye to replace the human crystalline lens in adult patients to improve vision (either distance vision or distance, intermediate, near vision; and additionally, optionally reduce the cylindrical power of the eye). | Medicontur hydrophilic posterior chamber trifocal IOLs are indicated to improve vision at far, intermediate and near distances in adults with cataract and/or ametropia (hyperopia, myopia) and/or presbyopia, secondary after removal of the crystalline lens. | P030102100202– IOLS, APHAKIC, MULTIFOCAL, ASPHERIC, HYDROPHILIC, ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Same as above | No | Yes | Moist heat | In-house | Yes | No | No | No | not available yet | Annex IX | N/A | Benz R&D, USA, Sarasota 6447 Parkland Drive, FL34230; Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
677MT | Trifocal Toric Bi-Flex Liberty | Posterior-chamber intraocular lens, pseudophakic, hydrophilic | 599302677MT2V | TF_01_B | Medicontur hydrophilic posterior chamber IOLs are intended for implantation into the capsular bag in the posterior chamber of the eye to replace the human crystalline lens in adult patients to improve vision (either distance vision or distance, intermediate, near vision; and additionally, optionally reduce the cylindrical power of the eye). | Medicontur hydrophilic posterior chamber trifocal toric IOLs are indicated to improve vision at far, intermediate and near distances and to reduce the cylindrical power of the eye in adults with cataract and/or ametropia (hyperopia, myopia, astigmatism), secondary to removal of the crystalline lens. | P030102100302– IOLs, APHAKIC, MULTIFOCAL, TORIC, HYDROPHILIC ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Same as above | No | Yes | Moist heat | In-house | Yes | No | No | No | not available yet | Annex IX | N/A | Benz R&D, USA, Sarasota 6447 Parkland Drive, FL34230; Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
677TA | Bi-Flex T | Posterior-chamber intraocular lens, pseudophakic, hydrophilic | 599302677TA2C | TF_01_B | Medicontur hydrophilic posterior chamber IOLs are intended for implantation into the capsular bag in the posterior chamber of the eye to replace the human crystalline lens in adult patients to improve vision (either distance vision or distance, intermediate, near vision; and additionally, optionally reduce the cylindrical power of the eye). | Medicontur hydrophilic posterior chamber monofocal toric IOLs are indicated to improve vision at far distance and to reduce the cylindrical power of the eye in adults with cataract and/or ametropia (hyperopia, myopia, astigmatism), secondary to removal of the crystalline lens. | P030102090302 – IOLS, APHAKIC, MONOFOCAL, TORIC, ASPHERIC, HYDROPHILIC, ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Same as above | No | Yes | Moist heat | In-house | Yes | No | No | No | not available yet | Annex IX | N/A | Benz R&D, USA, Sarasota 6447 Parkland Drive, FL34230; Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
690CM | Trifocal Z-Flex preloaded | Posterior-chamber intraocular lens, pseudophakic, hydrophilic | 599302690CMYR | TF_01_B | Medicontur hydrophilic posterior chamber IOLs are intended for implantation into the capsular bag in the posterior chamber of the eye to replace the human crystalline lens in adult patients to improve vision (either distance vision or distance, intermediate, near vision; and additionally, optionally reduce the cylindrical power of the eye). | Medicontur hydrophilic posterior chamber trifocal IOLs are indicated to improve vision at far, intermediate and near distances in adults with cataract and/or ametropia (hyperopia, myopia) and/or presbyopia, secondary after removal of the crystalline lens. | P030102100202– IOLS, APHAKIC, MULTIFOCAL, ASPHERIC, HYDROPHILIC, ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Same as above | No | Yes | Moist heat | In-house | Yes | No | No | Yes | not available yet | Annex IX | N/A | Benz R&D, USA, Sarasota 6447 Parkland Drive, FL34230; Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
690CTA | Z-Flex T Preloaded | Posterior-chamber intraocular lens, pseudophakic, hydrophilic | 599302690CTAQU | TF_01_B | Medicontur hydrophilic posterior chamber IOLs are intended for implantation into the capsular bag in the posterior chamber of the eye to replace the human crystalline lens in adult patients to improve vision (either distance vision or distance, intermediate, near vision; and additionally, optionally reduce the cylindrical power of the eye). | Medicontur hydrophilic posterior chamber monofocal toric IOLs are indicated to improve vision at far distance and to reduce the cylindrical power of the eye in adults with cataract and/or ametropia (hyperopia, myopia, astigmatism), secondary to removal of the crystalline lens. | P030102090302 – IOLS, APHAKIC, MONOFOCAL, TORIC, ASPHERIC, HYDROPHILIC, ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Same as above | No | Yes | Moist heat | In-house | Yes | No | No | Yes | not available yet | Annex IX | N/A | Benz R&D, USA, Sarasota 6447 Parkland Drive, FL34230; Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
690TA | Z-Flex T | Posterior-chamber intraocular lens, pseudophakic, hydrophilic | 599302690TAZL | TF_01_B | Medicontur hydrophilic posterior chamber IOLs are intended for implantation into the capsular bag in the posterior chamber of the eye to replace the human crystalline lens in adult patients to improve vision (either distance vision or distance, intermediate, near vision; and additionally, optionally reduce the cylindrical power of the eye). | Medicontur hydrophilic posterior chamber monofocal toric IOLs are indicated to improve vision at far distance and to reduce the cylindrical power of the eye in adults with cataract and/or ametropia (hyperopia, myopia, astigmatism), secondary to removal of the crystalline lens. | P030102090302 – IOLS, APHAKIC, MONOFOCAL, TORIC, ASPHERIC, HYDROPHILIC, ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Same as above | No | Yes | Moist heat | In-house | Yes | No | No | No | not available yet | Annex IX | N/A | Benz R&D, USA, Sarasota 6447 Parkland Drive, FL34230; Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
860PT | Z-Flex Toric POB-MA | Posterior chamber intraocular lens, pseudophakic, hydrophobic, preloaded in a single use injector | 599302860PT2J | TF_01_D | MEDICONTUR hydrophobic IOLs are intended for implantation into the capsular bag in the posterior chamber of the eye to replace the human crystalline lens in adult patients to improve distance vision or extend the depth of focus covering distance and intermediate vision. | Medicontur Toric IOLs are indicated for patients with corneal astigmatism who aspire to have improved uncorrected distance vision and reduction of residual refractive cylinder. | P030102090301 IOLs, APHAKIC, MONOFOCAL, TORIC, HYDROPHOBIC ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Hydrophobic IOLs are implantable devices and they: - are not intended to be placed in the teeth; - are not intended to be used in direct contact with the heart, the central circulatory system or the central nervous system; - do not have a biological effect or are not wholly or mainly absorbed; - are not intended to undergo chemical change in the body; - are not intended to administer medicinal products. - are not active implantable devices or their accessories; - are not breast implants or surgical meshes; - are not total or partial joint replacements; - are not spinal disc replacement implants, or, are not implantable devices that come into contact with the spinal column. | No | Yes | Ethylene oxide sterilisation (EtO) | In-house | Yes | No | No | No | not available yet | Annex IX | N/A | Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
860PTY | Z-Flex Toric POB-MA | Posterior chamber intraocular lens, pseudophakic, hydrophobic, blue light filter, preloaded in a single use injector | 599302860PTYU6 | TF_01_D | MEDICONTUR hydrophobic IOLs are intended for implantation into the capsular bag in the posterior chamber of the eye to replace the human crystalline lens in adult patients to improve distance vision or extend the depth of focus covering distance and intermediate vision. | Medicontur Toric IOLs are indicated for patients with corneal astigmatism who aspire to have improved uncorrected distance vision and reduction of residual refractive cylinder. | P030102090301 IOLs, APHAKIC, MONOFOCAL, TORIC, HYDROPHOBIC ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Same as above | No | Yes | Ethylene oxide sterilisation (EtO) | In-house | Yes | No | No | No | not available yet | Annex IX | N/A | Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
860PETY | Z-Flex EDOF Toric POB-MA | Posterior chamber intraocular lens, pseudophakic, hydrophobic, blue light filter, preloaded in a single use injector | 599302860PETYNW | TF_01_D | MEDICONTUR hydrophobic IOLs are intended for implantation into the capsular bag in the posterior chamber of the eye to replace the human crystalline lens in adult patients to improve distance vision or extend the depth of focus covering distance and intermediate vision. | EDOF toric models: Medicontur EDOF toric IOLs are indicated for patients with corneal astigmatism who aspire to have improved intermediate vision with uncompromised distance vision and reduction of residual refractive cylinder. | P030102090301 IOLs, APHAKIC, MONOFOCAL, TORIC, HYDROPHOBIC ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Same as above | No | Yes | Ethylene oxide sterilisation (EtO) | In-house | Yes | No | No | No | not available yet | Annex IX | N/A | Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
877PT | Bi-Flex Toric POB-MA | Posterior chamber intraocular lens, pseudophakic, hydrophobic, preloaded in a single use injector | 599302877PT3U | TF_01_D | MEDICONTUR hydrophobic IOLs are intended for implantation into the capsular bag in the posterior chamber of the eye to replace the human crystalline lens in adult patients to improve distance vision or extend the depth of focus covering distance and intermediate vision. | Medicontur Toric IOLs are indicated for patients with corneal astigmatism who aspire to have improved uncorrected distance vision and reduction of residual refractive cylinder. | P030102090301 IOLs, APHAKIC, MONOFOCAL, TORIC, HYDROPHOBIC ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Same as above | No | Yes | Ethylene oxide sterilisation (EtO) | In-house | Yes | No | No | No | not available yet | Annex IX | N/A | Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
877PTY | Bi-Flex Toric POB-MA | Posterior chamber intraocular lens, pseudophakic, hydrophobic, blue light filter, preloaded in a single use injector | 599302877PTYW2 | TF_01_D | MEDICONTUR hydrophobic IOLs are intended for implantation into the capsular bag in the posterior chamber of the eye to replace the human crystalline lens in adult patients to improve distance vision or extend the depth of focus covering distance and intermediate vision. | Medicontur Toric IOLs are indicated for patients with corneal astigmatism who aspire to have improved uncorrected distance vision and reduction of residual refractive cylinder. | P030102090301 IOLs, APHAKIC, MONOFOCAL, TORIC, HYDROPHOBIC ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Same as above | No | Yes | Ethylene oxide sterilisation (EtO) | In-house | Yes | No | No | No | not available yet | Annex IX | N/A | Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
877PETY | Bi-Flex ELON Toric POB-MA | Posterior chamber intraocular lens, pseudophakic, hydrophobic, blue light filter, preloaded in a single use injector | 599302877PETYRQ | TF_01_D | MEDICONTUR hydrophobic IOLs are intended for implantation into the capsular bag in the posterior chamber of the eye to replace the human crystalline lens in adult patients to improve distance vision or extend the depth of focus covering distance and intermediate vision. | EDOF toric models: Medicontur EDOF toric IOLs are indicated for patients with corneal astigmatism who aspire to have improved intermediate vision with uncompromised distance vision and reduction of residual refractive cylinder. | P030102090301 IOLs, APHAKIC, MONOFOCAL, TORIC, HYDROPHOBIC ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Same as above | No | Yes | Ethylene oxide sterilisation (EtO) | In-house | Yes | No | No | No | not available yet | Annex IX | N/A | Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
A45DT | 1stQ Add On | Posterior-chamber intraocular lens, pseudophakic, hydrophilic, for implantation into the ciliary sulcus | 599302A45DT4H | TF_01_E | MEDICONTUR Supplementary (Addon) IOLs are posterior chamber intraocular lenses intended for implantation into the ciliary sulcus in the posterior chamber of pseudophakic patients with a primary intraocular lens implanted in the capsular bag. | A45DT is indicated for pseudophakic patients with a primary IOL in the capsular bag who aspire to have improved near and intermediate vision with increased spectacle independence including a correction of corneal astigmatism and optionally refractive adjustment | P030102100302– IOLs, APHAKIC, MULTIFOCAL, TORIC, HYDROPHILIC ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Hydrophilic IOLs are implantable devices and they: - are not intended to be placed in the teeth; - are not intended to be used in direct contact with the heart, the central circulatory system or the central nervous system; - do not have a biological effect or are not wholly or mainly absorbed; - are not intended to undergo chemical change in the body; - are not intended to administer medicinal products. - are not active implantable devices or their accessories; - are not breast implants or surgical meshes; - are not total or partial joint replacements; - are not spinal disc replacement implants, or, are not implantable devices that come into contact with the spinal column. | No | Yes | Moist heat | In-house | Yes | No | No | No | not available yet | Annex IX | N/A | Benz R&D, USA, Sarasota 6447 Parkland Drive, FL34230; Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
A45RD2 | 1stQ Add On | Posterior-chamber intraocular lens, pseudophakic, hydrophilic, for implantation into the ciliary sulcus | 599302A45RD2U9 | TF_01_E | MEDICONTUR Supplementary (Addon) IOLs are posterior chamber intraocular lenses intended for implantation into the ciliary sulcus in the posterior chamber of pseudophakic patients with a primary intraocular lens implanted in the capsular bag. | A45RD2 is indicated for pseudophakic patients with a primary IOL in the capsular bag who aspire to have additional near and intermediate vision with increased spectacle independence and optionally a refractive adjustment. | P030102100202– IOLS, APHAKIC, MULTIFOCAL, ASPHERIC, HYDROPHILIC, ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Same as above | No | Yes | Moist heat | In-house | Yes | No | No | No | not available yet | Annex IX | N/A | Benz R&D, USA, Sarasota 6447 Parkland Drive, FL34230; Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
A45RT | 1stQ Add On | Posterior-chamber intraocular lens, pseudophakic, hydrophilic, for implantation into the ciliary sulcus | 599302A45RT5T | TF_01_E | MEDICONTUR Supplementary (Addon) IOLs are posterior chamber intraocular lenses intended for implantation into the ciliary sulcus in the posterior chamber of pseudophakic patients with a primary intraocular lens implanted in the capsular bag. | A45RT is indicated for pseudophakic patients with a primary IOL in the capsular bag who aspire to have a correction of their corneal astigmatism and optionally a refractive adjustment. | P030102090302 – IOLS, APHAKIC, MONOFOCAL, TORIC, ASPHERIC, HYDROPHILIC, ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Same as above | No | Yes | Moist heat | In-house | Yes | No | No | No | not available yet | Annex IX | N/A | Benz R&D, USA, Sarasota 6447 Parkland Drive, FL34230; Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
A45SML | Scharioth Macula Lens | Posterior-chamber intraocular lens, pseudophakic, hydrophilic, for implantation into the ciliary sulcus | 599302A45SMLWT | TF_01_E | MEDICONTUR Supplementary (Addon) IOLs are posterior chamber intraocular lenses intended for implantation into the ciliary sulcus in the posterior chamber of pseudophakic patients with a primary intraocular lens implanted in the capsular bag. | A45SML is indicated for pseudophakic patients with a dry form of age-related macular degeneration having a primary IOL in the capsular bag who aspire to have improved near vision. | P030102100202– IOLS, APHAKIC, MULTIFOCAL, ASPHERIC, HYDROPHILIC, ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Same as above | No | Yes | Moist heat | In-house | Yes | No | No | No | not available yet | Annex IX | N/A | Benz R&D, USA, Sarasota 6447 Parkland Drive, FL34230; Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
A46R | 1stQ Add On | Posterior-chamber intraocular lens, pseudophakic, hydrophilic, for implantation into the ciliary sulcus | 599302A46RFZ | TF_01_E | MEDICONTUR Supplementary (Addon) IOLs are posterior chamber intraocular lenses intended for implantation into the ciliary sulcus in the posterior chamber of pseudophakic patients with a primary intraocular lens implanted in the capsular bag. | A46R is indicated for pseudophakic patients with a primary IOL in the capsular bag who aspire to have a refractive adjustment | P030102090202 – IOLS, APHAKIC, MONOFOCAL, ASPHERIC, HYDROPHILIC, ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Same as above | No | Yes | Moist heat | In-house | Yes | No | No | No | not available yet | Annex IX | N/A | Benz R&D, USA, Sarasota 6447 Parkland Drive, FL34230; Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
A4EDF1 | 1stQ AddOn EDOF | Posterior-chamber intraocular lens, pseudophakic, hydrophilic, for implantation into the ciliary sulcus | 599302A4EDF1VG | TF_01_E | MEDICONTUR Supplementary (Addon) IOLs are posterior chamber intraocular lenses intended for implantation into the ciliary sulcus in the posterior chamber of pseudophakic patients with a primary intraocular lens implanted in the capsular bag. | A4EDF1, A4EDF2 are indicated for pseudophakic patients with a primary IOL in the capsular bag who aspire to have additional intermediate with increased spectacle independence and optionally a refractive adjustment. | P030102100202– IOLS, APHAKIC, MULTIFOCAL, ASPHERIC, HYDROPHILIC, ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Same as above | No | Yes | Moist heat | In-house | Yes | No | No | No | not available yet | Annex IX | N/A | Benz R&D, USA, Sarasota 6447 Parkland Drive, FL34230; Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
A4EDF2 | 1stQ AddOn EDOF | Posterior-chamber intraocular lens, pseudophakic, hydrophilic, for implantation into the ciliary sulcus | 599302A4EDF2VJ | TF_01_E | MEDICONTUR Supplementary (Addon) IOLs are posterior chamber intraocular lenses intended for implantation into the ciliary sulcus in the posterior chamber of pseudophakic patients with a primary intraocular lens implanted in the capsular bag. | A4EDF1, A4EDF2 are indicated for pseudophakic patients with a primary IOL in the capsular bag who aspire to have additional intermediate with increased spectacle independence and optionally a refractive adjustment. | P030102100202– IOLS, APHAKIC, MULTIFOCAL, ASPHERIC, HYDROPHILIC, ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Same as above | No | Yes | Moist heat | In-house | Yes | No | No | No | not available yet | Annex IX | N/A | Benz R&D, USA, Sarasota 6447 Parkland Drive, FL34230; Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
613AD | G-Flex | Posterior-chamber intraocular lens, pseudophakic, hydrophilic | 599302613ADWQ | TF_01_B | Medicontur hydrophilic posterior chamber IOLs are intended for implantation into the capsular bag in the posterior chamber of the eye to replace the human crystalline lens in adult patients to improve vision (either distance vision or distance, intermediate, near vision; and additionally, optionally reduce the cylindrical power of the eye). | Medicontur hydrophilic posterior chamber monofocal IOLs are indicated to improve vision at far distance in adults with cataract and/or ametropia (hyperopia, myopia), secondary to removal of the crystalline lens. | P030102090202 – IOLS, APHAKIC, MONOFOCAL, ASPHERIC, HYDROPHILIC, ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Same as above | No | Yes | Moist heat | In-house | Yes | No | No | No | not available yet | Annex IX | N/A | Benz R&D, USA, Sarasota 6447 Parkland Drive, FL34230; Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
860FAB | Z-Flex HB | Posterior chamber intraocular lens, pseudophakic, hydrophobic | 599302860FABPD | TF_01_C | MEDICONTUR hydrophobic IOLs are intended for implantation into the capsular bag in the posterior chamber of the eye to replace the human crystalline lens in adult patients to improve distance vision or extend the depth of focus covering distance and intermediate vision. | Medicontur hydrophobic monofocal IOLs are indicated to improve vision at far distance in aphakic adults with cataract and/or ametropia (hyperopia, myopia), secondary to the removal of the crystalline lens. | P030102090201 IOLs, APHAKIC, MONOFOCAL, ASPHERIC, HYDROPHOBIC ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Hydrophobic IOLs are implantable devices and they: - are not intended to be placed in the teeth; - are not intended to be used in direct contact with the heart, the central circulatory system or the central nervous system; - do not have a biological effect or are not wholly or mainly absorbed; - are not intended to undergo chemical change in the body; - are not intended to administer medicinal products. - are not active implantable devices or their accessories; - are not breast implants or surgical meshes; - are not total or partial joint replacements; - are not spinal disc replacement implants, or, are not implantable devices that come into contact with the spinal column. | No | Yes | Ethylene oxide sterilisation (EtO) | In-house | Yes | No | No | No | LB-002-5101-00-V04 POSTERIOR CHAMBER HB IOL IFU EN | Annex IX | N/A | Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
877FAB | Bi-Flex HB | Posterior chamber intraocular lens, pseudophakic, hydrophobic | 599302877FABR9 | TF_01_C | MEDICONTUR hydrophobic IOLs are intended for implantation into the capsular bag in the posterior chamber of the eye to replace the human crystalline lens in adult patients to improve distance vision or extend the depth of focus covering distance and intermediate vision. | Medicontur hydrophobic monofocal IOLs are indicated to improve vision at far distance in aphakic adults with cataract and/or ametropia (hyperopia, myopia), secondary to the removal of the crystalline lens. | P030102090201 IOLs, APHAKIC, MONOFOCAL, ASPHERIC, HYDROPHOBIC ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Same as above | No | Yes | Ethylene oxide sterilisation (EtO) | In-house | Yes | No | No | No | LB-002-5101-00-V04 POSTERIOR CHAMBER HB IOL IFU EN | Annex IX | N/A | Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
860FABY | Z-Flex HB | Posterior chamber intraocular lens, pseudophakic, hydrophobic, blue light filter | 599302860FABYJE | TF_01_C | MEDICONTUR hydrophobic IOLs are intended for implantation into the capsular bag in the posterior chamber of the eye to replace the human crystalline lens in adult patients to improve distance vision or extend the depth of focus covering distance and intermediate vision. | Medicontur hydrophobic monofocal IOLs are indicated to improve vision at far distance in aphakic adults with cataract and/or ametropia (hyperopia, myopia), secondary to the removal of the crystalline lens. | P030102090201 IOLs, APHAKIC, MONOFOCAL, ASPHERIC, HYDROPHOBIC ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Same as above | No | Yes | Ethylene oxide sterilisation (EtO) | In-house | Yes | No | No | No | LB-002-5101-00-V04 POSTERIOR CHAMBER HB IOL IFU EN | Annex IX | N/A | Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
877EBY | Bi-Flex ELON HB | Posterior chamber intraocular lens, pseudophakic, hydrophobic, blue light filter | 599302877EBYSM | TF_01_C | MEDICONTUR hydrophobic IOLs are intended for implantation into the capsular bag in the posterior chamber of the eye to replace the human crystalline lens in adult patients to improve distance vision or extend the depth of focus covering distance and intermediate vision. | Medicontur hydrophobic EDOF IOLs are indicated to improve vision at far and intermediate distances in aphakic adults with cataract and/or ametropia (hyperopia, myopia) and/or presbyopia, secondary to the removal of the crystalline lens. | P030102090201 IOLs, APHAKIC, MONOFOCAL, ASPHERIC, HYDROPHOBIC ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Same as above | No | Yes | Ethylene oxide sterilisation (EtO) | In-house | Yes | No | No | No | LB-002-5101-00-V04 POSTERIOR CHAMBER HB IOL IFU EN | Annex IX | N/A | Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
877FABY | Bi-Flex HB | Posterior chamber intraocular lens, pseudophakic, hydrophobic, blue light filter | 599302877FABYM8 | TF_01_C | MEDICONTUR hydrophobic IOLs are intended for implantation into the capsular bag in the posterior chamber of the eye to replace the human crystalline lens in adult patients to improve distance vision or extend the depth of focus covering distance and intermediate vision. | Medicontur hydrophobic monofocal IOLs are indicated to improve vision at far distance in aphakic adults with cataract and/or ametropia (hyperopia, myopia), secondary to the removal of the crystalline lens. | P030102090201 IOLs, APHAKIC, MONOFOCAL, ASPHERIC, HYDROPHOBIC ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Same as above | No | Yes | Ethylene oxide sterilisation (EtO) | In-house | Yes | No | No | No | LB-002-5101-00-V04 POSTERIOR CHAMBER HB IOL IFU EN | Annex IX | N/A | Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
860PA | Z-Flex POB-MA | Posterior chamber intraocular lens, pseudophakic, hydrophobic, preloaded in a single use injector | 599302860PAZ9 | TF_01_D | MEDICONTUR hydrophobic IOLs are intended for implantation into the capsular bag in the posterior chamber of the eye to replace the human crystalline lens in adult patients to improve distance vision or extend the depth of focus covering distance and intermediate vision. | Medicontur hydrophobic monofocal IOLs are indicated to improve vision at far distance in aphakic adults with cataract and/or ametropia (hyperopia, myopia), secondary to the removal of the crystalline lens. | P030102090201 IOLs, APHAKIC, MONOFOCAL, ASPHERIC, HYDROPHOBIC ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Same as above | No | Yes | Ethylene oxide sterilisation (EtO) | In-house | Yes | No | No | No | LB-006-5100-00-V11 POBMA Injector IFU EN | Annex IX | N/A | Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
877PA | Bi-Flex POB-MA | Posterior chamber intraocular lens, pseudophakic, hydrophobic, preloaded in a single use injector | 599302877PA2N | TF_01_D | MEDICONTUR hydrophobic IOLs are intended for implantation into the capsular bag in the posterior chamber of the eye to replace the human crystalline lens in adult patients to improve distance vision or extend the depth of focus covering distance and intermediate vision. | Medicontur hydrophobic monofocal IOLs are indicated to improve vision at far distance in aphakic adults with cataract and/or ametropia (hyperopia, myopia), secondary to the removal of the crystalline lens. | P030102090201 IOLs, APHAKIC, MONOFOCAL, ASPHERIC, HYDROPHOBIC ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Same as above | No | Yes | Ethylene oxide sterilisation (EtO) | In-house | Yes | No | No | No | LB-006-5100-00-V11 POBMA Injector IFU EN | Annex IX | N/A | Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
860PAY | Z-Flex POB-MA | Posterior chamber intraocular lens, pseudophakic, hydrophobic, blue light filter, preloaded in a single use injector | 599302860PAYSD | TF_01_D | MEDICONTUR hydrophobic IOLs are intended for implantation into the capsular bag in the posterior chamber of the eye to replace the human crystalline lens in adult patients to improve distance vision or extend the depth of focus covering distance and intermediate vision. | Medicontur hydrophobic monofocal IOLs are indicated to improve vision at far distance in aphakic adults with cataract and/or ametropia (hyperopia, myopia), secondary to the removal of the crystalline lens. | P030102090201 IOLs, APHAKIC, MONOFOCAL, ASPHERIC, HYDROPHOBIC ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Same as above | No | Yes | Ethylene oxide sterilisation (EtO) | In-house | Yes | No | No | No | LB-006-5100-00-V11 POBMA Injector IFU EN | Annex IX | N/A | Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
860PEY | Z-Flex EDOF POB-MA | Posterior chamber intraocular lens, pseudophakic, hydrophobic, blue light filter, preloaded in a single use injector | 599302860PEYSR | TF_01_D | MEDICONTUR hydrophobic IOLs are intended for implantation into the capsular bag in the posterior chamber of the eye to replace the human crystalline lens in adult patients to improve distance vision or extend the depth of focus covering distance and intermediate vision. | Medicontur hydrophobic EDOF IOLs are indicated to improve vision at far and intermediate distances in aphakic adults with cataract and/or ametropia (hyperopia, myopia) and/or presbyopia, secondary to the removal of the crystalline lens. | P030102090201 IOLs, APHAKIC, MONOFOCAL, ASPHERIC, HYDROPHOBIC ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Same as above | No | Yes | Ethylene oxide sterilisation (EtO) | In-house | Yes | No | No | No | LB-006-5100-00-V11 POBMA Injector IFU EN | Annex IX | N/A | Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
877PAY | Bi-Flex POB-MA | Posterior chamber intraocular lens, pseudophakic, hydrophobic, blue light filter, preloaded in a single use injector | 599302877PAYU9 | TF_01_D | MEDICONTUR hydrophobic IOLs are intended for implantation into the capsular bag in the posterior chamber of the eye to replace the human crystalline lens in adult patients to improve distance vision or extend the depth of focus covering distance and intermediate vision. | Medicontur hydrophobic monofocal IOLs are indicated to improve vision at far distance in aphakic adults with cataract and/or ametropia (hyperopia, myopia), secondary to the removal of the crystalline lens. | P030102090201 IOLs, APHAKIC, MONOFOCAL, ASPHERIC, HYDROPHOBIC ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Same as above | No | Yes | Ethylene oxide sterilisation (EtO) | In-house | Yes | No | No | No | LB-006-5100-00-V11 POBMA Injector IFU EN | Annex IX | N/A | Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
877PEY | Bi-Flex ELON POB-MA | Posterior chamber intraocular lens, pseudophakic, hydrophobic, blue light filter, preloaded in a single use injector | 599302877PEYUM | TF_01_D | MEDICONTUR hydrophobic IOLs are intended for implantation into the capsular bag in the posterior chamber of the eye to replace the human crystalline lens in adult patients to improve distance vision or extend the depth of focus covering distance and intermediate vision. | Medicontur hydrophobic EDOF IOLs are indicated to improve vision at far and intermediate distances in aphakic adults with cataract and/or ametropia (hyperopia, myopia) and/or presbyopia, secondary to the removal of the crystalline lens. | P030102090201 IOLs, APHAKIC, MONOFOCAL, ASPHERIC, HYDROPHOBIC ACRYLIC | 35658 | MDN1104 | MDS1005 | MDT2002, MDT2008, MDT2011 | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Same as above | No | Yes | Ethylene oxide sterilisation (EtO) | In-house | Yes | No | No | No | LB-006-5100-00-V11 POBMA Injector IFU EN | Annex IX | N/A | Medipack AG, Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | 3EC International a.s. Hranicna 18 Bratislava 82105 SLOVAKIA 2265
PIL-MA | Medjet PIL-MA injection system | Manual intraocular lens injector, single use | 599302MedjetPIL-MANP | TF_02 | Medicontur PIL-MA injector is intended to expel a prepared folded intraocular lens (IOL) from an IOL insertion cartridge into the posterior chamber of the eye during ophthalmic surgery. | Not applicable. | Q021111 | GMDN 47725 Manual intraocular lens injector, single use | MDN 1206 Non-active non-implantable ophthalmologic devices | MDS 1005 Devices in sterile condition MDS 1011 Devices in systems or procedure packs - SGS audit riportban ez nem szerepel! | MDT 2008 Devices manufactured in clean rooms and associated controlled environments MDT 2011 Devices which require packaging, including labelling MDT 2002 Devices manufactured using plastic processing | IIa | IIa | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 6 | PIL-MA injector is surgically invasive device intended for transient use and: - are not intended specifically to control, diagnose, monitor or correct a defect of the heart or of the central circulatory system through direct contact with those parts of the body, in which case they are classified as class III; - are not reusable surgical instruments, in which case they are classified as class I; - are not intended specifically for use in direct contact with the heart or central circulatory system or the central nervous system, in which case they are classified as class III; - are not intended to supply energy in the form of ionising radiation in which case they are classified as class IIb; - do not have a biological effect or are wholly or mainly absorbed in which case they are classified as class IIb; or - are not intended to administer medicinal products by means of a delivery system, if such administration of a medicinal product is done in a manner that is potentially hazardous taking account of the mode of application, in which case they are classified as class IIb. | No | Yes | Ethylene oxide sterilisation (EtO) | In-house | Yes | No | No | Yes | not available yet | Annex IX | N/A | Medipack AG Switzerland, Gewerbezentrum Arova Winterthurerstrasse, 8247 Flurlingen; AST Products Inc. USA, Billerica MA 01821, 9 Linnell Circle; Mikropakk Műanyag és Fémfeldolg. Kft. Hungary 1106 Budapest Jászberényi út 82. | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639
Medjet MC 1.6 | Medjet | Manual intraocular lens injector, single use | 599302MedjetMC1.67E | TF_03 | The Medicontur Medjet injectors are intended to be used by a trained ophthalmic surgeon for implantation of a foldable intraocular lens into the human eye. | Not applicable. | Q021111 | GMDN 47725 Manual intraocular lens injector, single use | MDN 1206 Non-active non-implantable ophthalmologic devices | MDS 1005 Devices in sterile condition | MDT 2008 Devices manufactured in clean rooms and associated controlled environments MDT 2011 Devices which require packaging, including labelling MDT 2002 Devices manufactured using plastic processing - SGS riportban szerepel, de ez biztosan kell? | IIa | IIa | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 6 | Medjet injectors are surgically invasive device intended for transient use and: - are not intended specifically to control, diagnose, monitor or correct a defect of the heart or of the central circulatory system through direct contact with those parts of the body, in which case they are classified as class III; - are not reusable surgical instruments, in which case they are classified as class I; - are not intended specifically for use in direct contact with the heart or central circulatory system or the central nervous system, in which case they are classified as class III; - are not intended to supply energy in the form of ionising radiation in which case they are classified as class IIb; - do not have a biological effect or are wholly or mainly absorbed in which case they are classified as class IIb; or - are not intended to administer medicinal products by means of a delivery system, if such administration of a medicinal product is done in a manner that is potentially hazardous taking account of the mode of application, in which case they are classified as class IIb. | No | Yes | Ethylene oxide sterilisation (EtO) | supplier | Yes | No | No | No | not available yet | Annex IX | Certificate number: FR19/81843446 | MDJ SAS Zone Industrielle Racine, Palladuc 63650 La Monnerie Le Montel, France | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639
Medjet MB 1.8 | Medjet | Manual intraocular lens injector, single use | 599302MedjetMB1.87B | TF_03 | The Medicontur Medjet injectors are intended to be used by a trained ophthalmic surgeon for implantation of a foldable intraocular lens into the human eye. | Not applicable. | Q021111 | GMDN 47725 Manual intraocular lens injector, single use | MDN 1206 Non-active non-implantable ophthalmologic devices | MDS 1005 Devices in sterile condition | MDT 2008 Devices manufactured in clean rooms and associated controlled environments MDT 2011 Devices which require packaging, including labelling MDT 2002 Devices manufactured using plastic processing - SGS riportban szerepel, de ez biztosan kell? | IIa | IIa | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 6 | Medjet injectors are surgically invasive device intended for transient use and: - are not intended specifically to control, diagnose, monitor or correct a defect of the heart or of the central circulatory system through direct contact with those parts of the body, in which case they are classified as class III; - are not reusable surgical instruments, in which case they are classified as class I; - are not intended specifically for use in direct contact with the heart or central circulatory system or the central nervous system, in which case they are classified as class III; - are not intended to supply energy in the form of ionising radiation in which case they are classified as class IIb; - do not have a biological effect or are wholly or mainly absorbed in which case they are classified as class IIb; or - are not intended to administer medicinal products by means of a delivery system, if such administration of a medicinal product is done in a manner that is potentially hazardous taking account of the mode of application, in which case they are classified as class IIb. | No | Yes | Ethylene oxide sterilisation (EtO) | supplier | Yes | No | No | No | not available yet | Annex IX | Certificate number: FR19/81843446 | MDJ SAS Zone Industrielle Racine, Palladuc 63650 La Monnerie Le Montel, France | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639
Medjet MA 2.2 | Medjet | Manual intraocular lens injector, single use | 599302MedjetMA2.26V | TF_03 | The Medicontur Medjet injectors are intended to be used by a trained ophthalmic surgeon for implantation of a foldable intraocular lens into the human eye. | Not applicable. | Q021111 | GMDN 47725 Manual intraocular lens injector, single use | MDN 1206 Non-active non-implantable ophthalmologic devices | MDS 1005 Devices in sterile condition | MDT 2008 Devices manufactured in clean rooms and associated controlled environments MDT 2011 Devices which require packaging, including labelling MDT 2002 Devices manufactured using plastic processing - SGS riportban szerepel, de ez biztosan kell? | IIa | IIa | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 6 | Medjet injectors are surgically invasive device intended for transient use and: - are not intended specifically to control, diagnose, monitor or correct a defect of the heart or of the central circulatory system through direct contact with those parts of the body, in which case they are classified as class III; - are not reusable surgical instruments, in which case they are classified as class I; - are not intended specifically for use in direct contact with the heart or central circulatory system or the central nervous system, in which case they are classified as class III; - are not intended to supply energy in the form of ionising radiation in which case they are classified as class IIb; - do not have a biological effect or are wholly or mainly absorbed in which case they are classified as class IIb; or - are not intended to administer medicinal products by means of a delivery system, if such administration of a medicinal product is done in a manner that is potentially hazardous taking account of the mode of application, in which case they are classified as class IIb. | No | Yes | Ethylene oxide sterilisation (EtO) | supplier | Yes | No | No | No | not available yet | Annex IX | Certificate number: FR19/81843446 | MDJ SAS Zone Industrielle Racine, Palladuc 63650 La Monnerie Le Montel, France | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639
Medjet MX 2.4 HB | Medjet | Manual intraocular lens injector, single use | 599302MedjetMX2.4HBVK | TF_03 | The Medicontur Medjet injectors are intended to be used by a trained ophthalmic surgeon for implantation of a foldable intraocular lens into the human eye. | Not applicable. | Q021111 | GMDN 47725 Manual intraocular lens injector, single use | MDN 1206 Non-active non-implantable ophthalmologic devices | MDS 1005 Devices in sterile condition | MDT 2008 Devices manufactured in clean rooms and associated controlled environments MDT 2011 Devices which require packaging, including labelling MDT 2002 Devices manufactured using plastic processing - SGS riportban szerepel, de ez biztosan kell? | IIa | IIa | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 6 | Medjet injectors are surgically invasive device intended for transient use and: - are not intended specifically to control, diagnose, monitor or correct a defect of the heart or of the central circulatory system through direct contact with those parts of the body, in which case they are classified as class III; - are not reusable surgical instruments, in which case they are classified as class I; - are not intended specifically for use in direct contact with the heart or central circulatory system or the central nervous system, in which case they are classified as class III; - are not intended to supply energy in the form of ionising radiation in which case they are classified as class IIb; - do not have a biological effect or are wholly or mainly absorbed in which case they are classified as class IIb; or - are not intended to administer medicinal products by means of a delivery system, if such administration of a medicinal product is done in a manner that is potentially hazardous taking account of the mode of application, in which case they are classified as class IIb. | No | Yes | Ethylene oxide sterilisation (EtO) | supplier | Yes | No | No | No | not available yet | Annex IX | Certificate number: FR19/81843446 | MDJ SAS Zone Industrielle Racine, Palladuc 63650 La Monnerie Le Montel, France | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639
Jetring 11ACB right | JetRing | Capsular Tension Ring Preloaded in a Single Use Injector | 599302Jetring11ACBrightEV | TF_07 | For the stabilization of the crystalline lens capsule in the presence of weak or partially absent zonules in patients undergoing cataract extraction with intraocular lens implantation. | The Medicontur Jetring CTR is indicated to expand and stabilize the capsular bag. | Q0299 | GMDN 42839 Sterile capsular tension ring preloaded in a single use injection kit | MDN 1104 Non-active soft tissue and other implants | MDS 1005 Devices in sterile condition | MDT 2008 evices manufactured in clean rooms and associated controlled environments MDT 2002 Devices manufactured using plastic processing - SGS riportban szerepel, de ez biztosan kell? MDT 2011 Devices which require packaging, including labelling | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Jetringd are implantable devices and they: - are not intended to be placed in the teeth; - are not intended to be used in direct contact with the heart, the central circulatory system or the central nervous system; - do not have a biological effect or are not wholly or mainly absorbed; - are not intended to undergo chemical change in the body; - are not intended to administer medicinal products. - are not active implantable devices or their accessories; - are not breast implants or surgical meshes; - are not total or partial joint replacements; - are not spinal disc replacement implants, or, are not implantable devices that come into contact with the spinal column. | No | Yes | Gamma irradiation | supplier | Yes | No | No | No | not available yet | Annex IX | Certificate number: G1 020543 0017 Rev.01 | Morcher GmbH Kapuzinerweg 12 70374 Stuttgart, Germany | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639
Jetring 12ACB right | JetRing | Capsular Tension Ring Preloaded in a Single Use Injector | 599302Jetring12ACBrightFL | TF_07 | For the stabilization of the crystalline lens capsule in the presence of weak or partially absent zonules in patients undergoing cataract extraction with intraocular lens implantation. | The Medicontur Jetring CTR is indicated to expand and stabilize the capsular bag. | Q0299 | GMDN 42839 Sterile capsular tension ring preloaded in a single use injection kit | MDN 1104 Non-active soft tissue and other implants | MDS 1005 Devices in sterile condition | MDT 2008 evices manufactured in clean rooms and associated controlled environments MDT 2002 Devices manufactured using plastic processing - SGS riportban szerepel, de ez biztosan kell? MDT 2011 Devices which require packaging, including labelling | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Jetringd are implantable devices and they: - are not intended to be placed in the teeth; - are not intended to be used in direct contact with the heart, the central circulatory system or the central nervous system; - do not have a biological effect or are not wholly or mainly absorbed; - are not intended to undergo chemical change in the body; - are not intended to administer medicinal products. - are not active implantable devices or their accessories; - are not breast implants or surgical meshes; - are not total or partial joint replacements; - are not spinal disc replacement implants, or, are not implantable devices that come into contact with the spinal column. | No | Yes | Gamma irradiation | supplier | Yes | No | No | No | not available yet | Annex IX | Certificate number: G1 020543 0017 Rev.01 | Morcher GmbH Kapuzinerweg 12 70374 Stuttgart, Germany | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639
VISCO-MC | N/A | Viscoelastic solution for intraocular use |  | TF_05 |  |  | Q02030301 | GMDN 35907 Vitreous/aqueous humour replacement medium, intraoperative | MDN 1206 Non-active non-implantable ophthalmologic devices | MDS 1005 Devices in sterile condition MDS 1011 Devices in systems or procedure packs | MDT 2008 Devices manufactured in clean rooms and associated controlled environments MDT 2011 Devices which require packaging, including labelling | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | Visco-MC products - are not intended to be placed in the teeth; - are not intended to be used in direct contact with the heart, the central circulatory system or the central nervous system; - do not have a biological effect or are not wholly or mainly absorbed; - are not intended to undergo chemical change in the body, except if the devices are placed in the teeth; - are not intended to administer medicinal products; - are not active implantable devices or their accessories; - are not breast implants or surgical meshes; - are not total or partial joint replacements, with the exception of ancillary components such as screws, wedges, plates and instruments; or - are not spinal disc replacement implants nor implantable devices that come into contact with the spinal column, with the exception of components such as screws, wedges, plates and instruments. | No | Yes | Ethylene oxide sterilisation (EtO) Moist heat (steam) | supplier | Yes |  | No |  | not available yet | Annex IX | Certificate number: G1 027023 0068 Rev. 01 | Croma Pharma Gmbh Industriezeile 6. 2100 - Loebendorf AT - Austria | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639
BioVis 1.6% | N/A | Viscoelastic solution for intraocular use | 599302BioVisXL | TF_04 | The intended use/purpose and application of the device include facilitating the intraocular surgery by protecting the tissue of the eye from the physical trauma and maintaining the shape of the eye during the surgical procedure. | Intraocular surgery of the anterior segment, including cataract extraction and intraocular lens implantation. The viscoelastic properties of BioVis cohesive products allow lubrication, support and protection of ocular tissues during ophthalmic surgery. Maintains the depth of the anterior chamber during surgery, allows for efficient manipulation with reduced trauma to the corneal endothelium and other surrounding tissues. | Q02030301 | GMDN 35907 Vitreous/aqueous humour replacement medium, intraoperative | MDN 1206 Non-active non-implantable ophthalmologic devices | MDS 1005 Devices in sterile condition MDS 1011 Devices in systems or procedure packs | MDT 2008 Devices manufactured in clean rooms and associated controlled environments MDT 2011 Devices which require packaging, including labelling | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | BioVis products - are not intended to be placed in the teeth; - are not intended to be used in direct contact with the heart, the central circulatory system or the central nervous system; - do not have a biological effect or are not wholly or mainly absorbed; - are not intended to undergo chemical change in the body, except if the devices are placed in the teeth; - are not intended to administer medicinal products; - are not active implantable devices or their accessories; - are not breast implants or surgical meshes; - are not total or partial joint replacements, with the exception of ancillary components such as screws, wedges, plates and instruments; or - are not spinal disc replacement implants nor implantable devices that come into contact with the spinal column, with the exception of components such as screws, wedges, plates and instruments. | No | Yes | Ethylene oxide sterilisation (EtO) Moist heat (steam) | supplier | Yes |  | No |  | not available yet | Annex IX | Certificate number: 1984-MDD-18-485 | VSY Biyoteknoloji ve İlaç Sanayi A.Ş. İstanbul Tuzla Organize Sanayi Bölgesi 3. Cad. No: 3 Tepeören Tuzla-İstanbul, Turkey | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639
BioVis 1.8% | N/A | Viscoelastic solution for intraocular use | 599302BioVisXL | TF_04 | The intended use/purpose and application of the device include facilitating the intraocular surgery by protecting the tissue of the eye from the physical trauma and maintaining the shape of the eye during the surgical procedure. | Intraocular surgery of the anterior segment, including cataract extraction and intraocular lens implantation. The viscoelastic properties of BioVis cohesive products allow lubrication, support and protection of ocular tissues during ophthalmic surgery. Maintains the depth of the anterior chamber during surgery, allows for efficient manipulation with reduced trauma to the corneal endothelium and other surrounding tissues. | Q02030301 | GMDN 35907 Vitreous/aqueous humour replacement medium, intraoperative | MDN 1206 Non-active non-implantable ophthalmologic devices | MDS 1005 Devices in sterile condition MDS 1011 Devices in systems or procedure packs | MDT 2008 Devices manufactured in clean rooms and associated controlled environments MDT 2011 Devices which require packaging, including labelling | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | BioVis products - are not intended to be placed in the teeth; - are not intended to be used in direct contact with the heart, the central circulatory system or the central nervous system; - do not have a biological effect or are not wholly or mainly absorbed; - are not intended to undergo chemical change in the body, except if the devices are placed in the teeth; - are not intended to administer medicinal products; - are not active implantable devices or their accessories; - are not breast implants or surgical meshes; - are not total or partial joint replacements, with the exception of ancillary components such as screws, wedges, plates and instruments; or - are not spinal disc replacement implants nor implantable devices that come into contact with the spinal column, with the exception of components such as screws, wedges, plates and instruments. | No | Yes | Ethylene oxide sterilisation (EtO) Moist heat (steam) | supplier | Yes |  | No |  | not available yet | Annex IX | Certificate number: 1984-MDD-18-485 | VSY Biyoteknoloji ve İlaç Sanayi A.Ş. İstanbul Tuzla Organize Sanayi Bölgesi 3. Cad. No: 3 Tepeören Tuzla-İstanbul, Turkey | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639
BioVis 3.0% | N/A | Viscoelastic solution for intraocular use | 599302BioVisXL | TF_04 | The intended use/purpose and application of the device include facilitating the intraocular surgery by protecting the tissue of the eye from the physical trauma and maintaining the shape of the eye during the surgical procedure. | Intraocular surgery of the anterior segment, including cataract extraction and intraocular lens implantation. The viscoelastic properties of BioVis cohesive products allow lubrication, support and protection of ocular tissues during ophthalmic surgery. Maintains the depth of the anterior chamber during surgery, allows for efficient manipulation with reduced trauma to the corneal endothelium and other surrounding tissues. | Q02030301 | GMDN 35907 Vitreous/aqueous humour replacement medium, intraoperative | MDN 1206 Non-active non-implantable ophthalmologic devices | MDS 1005 Devices in sterile condition MDS 1011 Devices in systems or procedure packs | MDT 2008 Devices manufactured in clean rooms and associated controlled environments MDT 2011 Devices which require packaging, including labelling | IIb | IIb | Annex VIII of Regulation (EU) 2017/745 (MDR); rule 8 | BioVis products - are not intended to be placed in the teeth; - are not intended to be used in direct contact with the heart, the central circulatory system or the central nervous system; - do not have a biological effect or are not wholly or mainly absorbed; - are not intended to undergo chemical change in the body, except if the devices are placed in the teeth; - are not intended to administer medicinal products; - are not active implantable devices or their accessories; - are not breast implants or surgical meshes; - are not total or partial joint replacements, with the exception of ancillary components such as screws, wedges, plates and instruments; or - are not spinal disc replacement implants nor implantable devices that come into contact with the spinal column, with the exception of components such as screws, wedges, plates and instruments. | No | Yes | Ethylene oxide sterilisation (EtO) Moist heat (steam) | supplier | Yes |  | No |  | not available yet | Annex IX | Certificate number: 1984-MDD-18-485 | VSY Biyoteknoloji ve İlaç Sanayi A.Ş. İstanbul Tuzla Organize Sanayi Bölgesi 3. Cad. No: 3 Tepeören Tuzla-İstanbul, Turkey | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639 | SGS Belgium NV SGS House Noorderlaan 87 2030 Antwerp Belgium 1639

## Listák
TF
TF_01_B
TF_01_C
TF_01_D
TF_01_E
TF_02
TF_03
TF_04_A
TF_04_B
TF_05
TF_07